How To Use DecOcx.ocx

The new version PCI12BITADDA Jumpers:

============================================================

JP1 – Programmable Gain Control Factor（AD526）

Short – No AD526(Default)

Open – Has AD526

JP2 – Input Mode for LF398

The JP2 is used to adjust input offset of LF398. Setting the jumper to pin 2 and pin 3 will ground the input value（0V）, so the data that will be converted should also be 0. This setting is used whenever we need to calibrate the VR1 variable resistors so that the minimum voltage will be attained.

Shorting jumper to pin 1 and pin 2 will make use of the channel inputs to be read and converted. It is the normal input mode.

Short 2,3 – Ground

Short 1,2 – Normal Input

JP3 – Input Mode for SPT574

The JP3 is used to adjust input offset of SPT574. Setting the jumper to pin 2 and pin 3 will ground the input value（0V）, so the data that will be converted should also be 0. This setting is used whenever we need to calibrate the VR2 and VR4 variable resistors so that the minimum voltage will be attained.

Shorting jumper to pin 1 and pin 2 will make use of the channel inputs to be read and converted. It is the normal input mode.

Short 2,3 – Ground

Short 1,2 – Normal Input

JP4 – Voltage Range

Short 1,2 – 10 Volts

Short 2,3 – 20 Volts

JP5/JP6/JP7 – Unipolar/Bipolar

| JP5 | JP6/JP7 | Polarity |
| --- | --- | --- |
| Short 2,3 | Short 3,4 | unipolar |
| Short 1,2 | Short 1,2 | bipolar |

The following methods are exposed by the OCX control:

============================================================

Methods

Closed

Type

None

Parameters

None

Example(Visual Basic)

Dim adr As Long

If Not DecOcx1.GetDevicePortAddress(adr) Then

DecOcx1.Closed		// Close the device.

End

End If

Methods

GetDevicePortAddress

Type

Boolean

ReadOnly

Parameters

Address as Long

Example(Visual Basic)

Dim adr As Long

If Not DecOcx1.GetDevicePortAddress(adr) Then

DecOcx1.Closed		// Close the device.

End

Else

Msgbox “The I/O Address is” + adr

End If

Methods

GetPortByte

Type

Integer

ReadOnly

Parameters

PortAddress as Long

State as Byte

Example(Visual Basic)

Dim adr As Long

Dim state As Byte

Dim DI as Integer

If Not DecOcx1.GetDevicePortAddress(adr) Then

DecOcx1.Closed		// Close the device.

End

Else

// To get state of first byte of port+6

DI = DecOcx1.GetPortByte(adr + 6, 0)

Msgbox “The first byte of Port+6 is ” + DI

End If

Methods

GetPortLong

Type

Long

ReadOnly

Parameters

PortAddress as Long

Example(Visual Basic)

Dim adr As Long

Dim DI as Long

If Not DecOcx1.GetDevicePortAddress(adr) Then

DecOcx1.Closed		// Close the device.

End

Else

‘’ To get state of port+6

DI = DecOcx1.GetPortLong(adr + 6)

Msgbox “The Port+6 is” + DI

End If

Methods

SetPortByte

Type

Boolean

Parameters

PortAddress as Long

State as Byte

Example(Visual Basic)

Dim adr As Long

Dim state As Byte

If Not DecOcx1.GetDevicePortAddress(adr) Then

DecOcx1.Closed		// Close the device.

End

Else

‘’ To set the first byte of port+6 to be state

State = 128

If DecOcx1.SetPortByte(adr + 6, state) Then

Msgbox “Set the first byte of Port+6 to be “ + state

End If

Methods

SetPortLong

Type

Boolean

Parameters

PortAddress as Long

State as Long

Example(Visual Basic)

Dim adr As Long

Dim state As Long

If Not DecOcx1.GetDevicePortAddress(adr) Then

DecOcx1.Closed		// Close the device.

End

Else

‘’ To set the first byte of port+6 to be state

state = 4095

If DecOcx1.SetPortByte(adr + 6, state) Then

Msgbox “Set the Port+6 to be “ + state

End If